## Gain DIP Switch Configuration
FUTEK Advanced Sensor Technology, Inc.
IAA100 Gain and Shunt Settings
1. Excitation | 10 | VDC | Directions: 1. Select excitation voltage to sensor from amplifier. 5V or 10V can be set on the amplifier board. 2. Enter the sensitivity (mV/V) of the sensor as found on sensor calibration certificate. 3. Choose amplifier full output level of +/- 5VDC or +/- 10VDC. 4. The yellow rows in the table below reflect gain DIP switch combinations that refelect a gain multiple to amplify the sensors output to the select amplifier full output.
2. Sensor's Full Output | 2 | mV/V
3. Amplifier Full Output | 10 | VDC
Switch Down | 0 | Off
Switch Up | 1 | On
Select the highlighted row that will provide the gain that results in the output voltage closest to the specified value.
Vout = Excitation * Sensor mV/V * Gain
Below are DIP switch configurations that will satisfy the above input information.
SW3 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Sensitivity (mV/V) | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 10
Resistance (Ohms) | 23.7 | 47.5 | 71.5 | 95.3 | 124 | 143 | 191 | 475 | Combined Resistance | Gain | Output (VDC)
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 1 | 0.02
1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 475 | 106.2632 | 2.1253
2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 191 | 262.7801 | 5.2556
3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 143 | 350.6503 | 7.013
4 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 136.2237 | 368.0433 | 7.3609
5 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 124 | 404.2258 | 8.0845
6 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 109.911 | 455.9135 | 9.1183
7 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 98.3306 | 509.489 | 10.1898
8 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 95.3 | 525.659 | 10.5132
9 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 81.7754 | 612.4305 | 12.2486
10 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 79.3749 | 630.9221 | 12.6184
11 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 75.1873 | 666.0059 | 13.3201
12 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 71.5 | 700.3007 | 14.006
13 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 69.7648 | 717.6936 | 14.3539
14 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 66.412 | 753.8762 | 15.0775
15 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 64.9124 | 771.2691 | 15.4254
16 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 63.5777 | 787.4391 | 15.7488
17 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 62.1455 | 805.5639 | 16.1113
18 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 58.2656 | 859.1393 | 17.1828
19 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 57.188 | 875.3093 | 17.5062
20 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 56.0725 | 892.7022 | 17.854
21 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 53.886 | 928.8848 | 18.5777
22 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 52.0248 | 962.0808 | 19.2416
23 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 51.0427 | 980.5725 | 19.6114
24 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 49.2778 | 1015.6563 | 20.3131
25 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 48.3958 | 1034.1479 | 20.683
26 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 47.6667 | 1049.951 | 20.999
27 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 47.5 | 1053.6316 | 21.0726
28 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 46.8892 | 1067.344 | 21.3469
29 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 45.3504 | 1103.5265 | 22.0705
30 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 44.6461 | 1120.9194 | 22.4184
31 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 44.0106 | 1137.0894 | 22.7418
32 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 43.3195 | 1155.2142 | 23.1043
33 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 43.1818 | 1158.8947 | 23.1779
34 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 42.0286 | 1190.6649 | 23.8133
35 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 41.3979 | 1208.7897 | 24.1758
36 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 40.851 | 1224.9597 | 24.4992
37 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 40.2786 | 1242.3526 | 24.8471
38 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 39.1379 | 1278.5351 | 25.5707
39 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 38.6122 | 1295.928 | 25.9186
40 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 38.1466 | 1311.7312 | 26.2346
41 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 38.0398 | 1315.4117 | 26.3082
42 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 37.616 | 1330.2228 | 26.6045
43 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 36.6487 | 1365.3066 | 27.3061
44 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 36.1586 | 1383.7983 | 27.676
45 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 35.6562 | 1403.2819 | 28.0656
46 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 35.3109 | 1416.9943 | 28.3399
47 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 35.2193 | 1420.6748 | 28.4135
48 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 34.4311 | 1453.1769 | 29.0635
49 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 34.344 | 1456.8574 | 29.1371
50 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 34.0236 | 1470.5698 | 29.4114
51 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 33.6533 | 1486.7398 | 29.7348
52 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 33.1665 | 1508.5451 | 30.1709
53 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 32.482 | 1540.3152 | 30.8063
54 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 32.104 | 1558.44 | 31.1688
55 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 32.0283 | 1562.1205 | 31.2424
56 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 31.7741 | 1574.61 | 31.4922
57 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 31.6999 | 1578.2906 | 31.5658
58 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 31.4267 | 1592.0029 | 31.8401
59 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 30.7279 | 1628.1855 | 32.5637
60 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 30.4029 | 1645.5784 | 32.9116
61 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 30.047 | 1665.062 | 33.3012
62 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 29.7819 | 1679.8732 | 33.5975
63 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 29.7167 | 1683.5537 | 33.6711
64 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 29.1723 | 1714.957 | 34.2991
65 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 29.1098 | 1718.6375 | 34.3727
66 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 28.8609 | 1733.4486 | 34.669
67 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 28.5399 | 1752.9323 | 35.0586
68 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 28.2594 | 1770.3252 | 35.4065
69 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27.693 | 1806.5077 | 36.1302
70 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 27.4843 | 1820.2201 | 36.4044
71 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 27.4288 | 1823.9006 | 36.478
72 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 27.2422 | 1836.3901 | 36.7278
73 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 27.1876 | 1840.0707 | 36.8014
74 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 26.9223 | 1858.1954 | 37.1639
75 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 26.4695 | 1889.9656 | 37.7993
76 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 26.1674 | 1911.7709 | 38.2354
77 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 25.9479 | 1927.9409 | 38.5588
78 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 25.7645 | 1941.6533 | 38.8331
79 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 25.7157 | 1945.3338 | 38.9067
80 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 25.2929 | 1977.8358 | 39.5567
81 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 25.2459 | 1981.5164 | 39.6303
82 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 25.0723 | 1995.2287 | 39.9046
83 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 24.8298 | 2014.7124 | 40.2942
84 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 24.6038 | 2033.2041 | 40.6641
85 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 24.1863 | 2068.2878 | 41.3658
86 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 24.0142 | 2083.099 | 41.662
87 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 23.9719 | 2086.7795 | 41.7356
88 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 23.7916 | 2102.5826 | 42.0517
89 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 23.7 | 2110.7046 | 42.2141
90 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 23.5963 | 2119.9755 | 42.3995
91 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 23.2002 | 2156.1581 | 43.1232
92 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 23.0144 | 2173.551 | 43.471
93 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 22.8444 | 2189.721 | 43.7944
94 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 22.6568 | 2207.8458 | 44.1569
95 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 22.5737 | 2215.9678 | 44.3194
96 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 22.3352 | 2239.6159 | 44.7923
97 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 22.2986 | 2243.2965 | 44.8659
98 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 22.1198 | 2261.4212 | 45.2284
99 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 21.9627 | 2277.5912 | 45.5518
100 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 21.7961 | 2294.9842 | 45.8997
101 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 21.4577 | 2331.1667 | 46.6233
102 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 21.3322 | 2344.8791 | 46.8976
103 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 21.2987 | 2348.5596 | 46.9712
104 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 21.1563 | 2364.3627 | 47.2873
105 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 21.0838 | 2372.4847 | 47.4497
106 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 20.992 | 2382.8544 | 47.6571
107 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 20.6873 | 2417.9382 | 48.3588
108 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 20.5303 | 2436.4299 | 48.7286
109 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 20.3305 | 2460.355 | 49.2071
110 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 20.2542 | 2469.6259 | 49.3925
111 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 20.1878 | 2477.7479 | 49.555
112 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 19.9616 | 2505.8084 | 50.1162
113 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 19.8971 | 2513.9304 | 50.2786
114 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 19.824 | 2523.2013 | 50.464
115 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 19.6977 | 2539.3714 | 50.7874
116 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 19.4961 | 2565.6181 | 51.3124
117 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 19.2905 | 2592.9468 | 51.8589
118 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 19.1566 | 2611.0716 | 52.2214
119 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 19.0971 | 2619.1936 | 52.3839
120 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 19.0386 | 2627.2416 | 52.5448
121 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 18.9799 | 2635.3636 | 52.7073
122 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 18.9134 | 2644.6345 | 52.8927
123 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 18.658 | 2680.8171 | 53.6163
124 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 18.5377 | 2698.21 | 53.9642
125 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 18.3747 | 2722.1351 | 54.4427
126 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 18.3049 | 2732.5048 | 54.6501
127 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 18.2507 | 2740.6268 | 54.8125
128 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 18.0728 | 2767.5885 | 55.3518
129 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 18.0199 | 2775.7106 | 55.5142
130 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 17.9528 | 2786.0802 | 55.7216
131 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 17.7999 | 2810.0053 | 56.2001
132 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 17.6904 | 2827.3983 | 56.548
133 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 17.4668 | 2863.5808 | 57.2716
134 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 17.4104 | 2872.8517 | 57.457
135 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 17.3613 | 2880.9737 | 57.6195
136 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 17.3129 | 2889.0217 | 57.7804
137 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 17.2643 | 2897.1437 | 57.9429
138 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 17.157 | 2915.2685 | 58.3054
139 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 16.9976 | 2942.5972 | 58.8519
140 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 16.8472 | 2968.844 | 59.3769
141 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 16.756 | 2985.014 | 59.7003
142 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 16.7041 | 2994.2849 | 59.8857
143 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 16.6589 | 3002.4069 | 60.0481
144 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 16.5046 | 3030.4674 | 60.6093
145 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 16.4604 | 3038.5894 | 60.7718
146 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 16.4103 | 3047.8603 | 60.9572
147 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 16.2825 | 3071.7854 | 61.4357
148 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 16.185 | 3090.2771 | 61.8055
149 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 16.0033 | 3125.3609 | 62.5072
150 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 15.9503 | 3135.7306 | 62.7146
151 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 15.9091 | 3143.8526 | 62.8771
152 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 15.8295 | 3159.6557 | 63.1931
153 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 15.8111 | 3163.3362 | 63.2667
154 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 15.7428 | 3177.0486 | 63.541
155 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 15.5655 | 3213.2311 | 64.2646
156 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 15.4817 | 3230.6241 | 64.6125
157 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 15.4046 | 3246.7941 | 64.9359
158 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 15.319 | 3264.9188 | 65.2984
159 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 15.3018 | 3268.5994 | 65.372
160 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 15.1918 | 3292.2475 | 65.845
161 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 15.1544 | 3300.3695 | 66.0074
162 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 15.0716 | 3318.4943 | 66.3699
163 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 14.9985 | 3334.6643 | 66.6933
164 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 14.9207 | 3352.0572 | 67.0411
165 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 14.7613 | 3388.2398 | 67.7648
166 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 14.721 | 3397.5107 | 67.9502
167 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 14.6859 | 3405.6327 | 68.1127
168 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 14.618 | 3421.4358 | 68.4287
169 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 14.6023 | 3425.1163 | 68.5023
170 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 14.5394 | 3439.9275 | 68.7985
171 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 14.3926 | 3475.0113 | 69.5002
172 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 14.3164 | 3493.5029 | 69.8701
173 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 14.237 | 3512.9866 | 70.2597
174 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 14.1816 | 3526.699 | 70.534
175 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 14.1668 | 3530.3795 | 70.6076
176 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 14.0375 | 3562.8815 | 71.2576
177 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 14.023 | 3566.562 | 71.3312
178 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 13.9693 | 3580.2744 | 71.6055
179 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 13.9065 | 3596.4444 | 71.9289
180 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 13.8227 | 3618.2497 | 72.365
181 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 13.7023 | 3650.0199 | 73.0004
182 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 13.6346 | 3668.1447 | 73.3629
183 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 13.6209 | 3671.8252 | 73.4365
184 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 13.5747 | 3684.3147 | 73.6863
185 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 13.5612 | 3687.9952 | 73.7599
186 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 13.5109 | 3701.7076 | 74.0342
187 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 13.3801 | 3737.8901 | 74.7578
188 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 13.3181 | 3755.283 | 75.1057
189 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 13.2494 | 3774.7667 | 75.4953
190 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 13.1976 | 3789.5778 | 75.7916
191 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 13.1848 | 3793.2583 | 75.8652
192 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 13.0765 | 3824.6616 | 76.4932
193 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 13.0639 | 3828.3421 | 76.5668
194 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 13.0135 | 3843.1533 | 76.8631
195 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 12.9479 | 3862.6369 | 77.2527
196 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 12.8898 | 3880.0298 | 77.6006
197 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 12.7707 | 3916.2124 | 78.3242
198 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 12.7261 | 3929.9248 | 78.5985
199 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 12.7142 | 3933.6053 | 78.6721
200 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 12.674 | 3946.0948 | 78.9219
201 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 12.6622 | 3949.7753 | 78.9955
202 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 12.6043 | 3967.9001 | 79.358
203 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 12.5042 | 3999.6702 | 79.9934
204 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 12.4363 | 4021.4755 | 80.4295
205 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 12.3865 | 4037.6455 | 80.7529
206 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 12.3446 | 4051.3579 | 81.0272
207 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 12.3334 | 4055.0385 | 81.1008
208 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 12.2353 | 4087.5405 | 81.7508
209 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 12.2243 | 4091.221 | 81.8244
210 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 12.1834 | 4104.9334 | 82.0987
211 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 12.1259 | 4124.417 | 82.4883
212 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 12.0717 | 4142.9087 | 82.8582
213 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 11.9703 | 4177.9925 | 83.5598
214 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 11.928 | 4192.8036 | 83.8561
215 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 11.9176 | 4196.4842 | 83.9297
216 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 11.8729 | 4212.2873 | 84.2457
217 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 11.824 | 4229.6802 | 84.5936
218 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 11.7237 | 4265.8627 | 85.3173
219 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 11.6761 | 4283.2556 | 85.6651
220 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 11.6322 | 4299.4256 | 85.9885
221 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 11.5833 | 4317.5504 | 86.351
222 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 11.4987 | 4349.3206 | 86.9864
223 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 11.489 | 4353.0011 | 87.06
224 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 11.4413 | 4371.1259 | 87.4225
225 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 11.3991 | 4387.2959 | 87.7459
226 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 11.3541 | 4404.6888 | 88.0938
227 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 11.2616 | 4440.8713 | 88.8174
228 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 11.2269 | 4454.5837 | 89.0917
229 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 11.2176 | 4458.2643 | 89.1653
230 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 11.178 | 4474.0674 | 89.4813
231 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 11.132 | 4492.559 | 89.8512
232 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 11.0457 | 4527.6428 | 90.5529
233 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 11.0008 | 4546.1345 | 90.9227
234 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 10.921 | 4579.3305 | 91.5866
235 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 10.8354 | 4615.5131 | 92.3103
236 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 10.7947 | 4632.906 | 92.6581
237 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 10.7571 | 4649.076 | 92.9815
238 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 10.6346 | 4702.6515 | 94.053
239 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 10.5937 | 4720.7762 | 94.4155
240 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 10.5576 | 4736.9462 | 94.7389
241 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 10.5189 | 4754.3392 | 95.0868
242 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 10.4395 | 4790.5217 | 95.8104
243 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 10.4017 | 4807.9146 | 96.1583
244 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 10.328 | 4842.2094 | 96.8442
245 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 10.2537 | 4877.2932 | 97.5459
246 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 10.215 | 4895.7849 | 97.9157
247 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 10.037 | 4982.5563 | 99.6511
248 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 10.0045 | 4998.7263 | 99.9745
249 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 9.8984 | 5052.3018 | 101.046
250 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 9.7982 | 5103.9895 | 102.0798
251 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 9.7292 | 5140.172 | 102.8034
252 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 9.6964 | 5157.565 | 103.1513
253 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 9.5339 | 5245.4352 | 104.9087
254 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 9.2576 | 5401.9522 | 108.039
255 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 9.0806 | 5507.2153 | 110.1443

## Shunt DIP Switch Configuration
FUTEK Advanced Sensor Technology, Inc.
IAA100 Gain and Shunt Settings
1. Excitation | 10 | VDC | Directions: 1. Select excitation voltage to sensor from amplifier. 5V or 10V can be set on the amplifier board. 2. Enter the bridge resistance (Ω) of the sensor between the +/- Signal connections. 3. Enter the gain of the system. The Gain DIP Switch Configuration worksheet in Column K shows the gain level for different gain DIP switch settings. 4.Choose amplifier full output level.
2. Bridge Resistance | 350 | Ω
3. Gain | 509
4. Amplifier Full Output | 10 | VDC
Switch Down | 0 | Off
Switch Up | 1 | On
Select the highlighted row that will provide the shunt value that results in the output voltage closest to the specified value.
Below are DIP switch configurations that will satisfy the above input information.
SW5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Resistance (Ohms) | 30000 | 43700 | 60400 | 87600 | 100000 | 150000 | 300000 | 432000 | Combined Resistance | RT | V1-V2 | Output (VDC)
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 432000 | 349.7167 | 0.0020246 | 1.0305
2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 300000 | 349.5921 | 0.002915 | 1.4837
3 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 177049.1803 | 349.3095 | 0.0049372 | 2.5131
4 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 150000 | 349.1852 | 0.0058265 | 2.9657
5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 111340.2062 | 348.9032 | 0.0078465 | 3.9938
6 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 100000 | 348.7793 | 0.0087347 | 4.446
7 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 100000 | 348.7793 | 0.0087347 | 4.446
8 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 87600 | 348.6072 | 0.0099687 | 5.0741
9 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 81203.0075 | 348.4979 | 0.010752 | 5.4729
10 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 81203.0075 | 348.4979 | 0.010752 | 5.4729
11 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 75000 | 348.3743 | 0.01164 | 5.9245
12 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 72831.4088 | 348.3261 | 0.011985 | 6.1005
13 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 67801.8576 | 348.2025 | 0.012872 | 6.5519
14 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 63905.3254 | 348.0935 | 0.013655 | 6.9503
15 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 60400 | 347.9835 | 0.014445 | 7.3525
16 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 60000 | 347.9702 | 0.014541 | 7.4013
17 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 58604.0288 | 347.9221 | 0.014886 | 7.5771
18 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 55303.0303 | 347.7989 | 0.015772 | 8.028
19 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 52991.0642 | 347.7035 | 0.016458 | 8.3771
20 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 52682.9268 | 347.6901 | 0.016554 | 8.4259
21 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 50277.4695 | 347.5804 | 0.017343 | 8.8276
22 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 50000 | 347.567 | 0.017439 | 8.8764
23 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 49026.8018 | 347.5191 | 0.017784 | 9.052
24 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 46695.0959 | 347.3961 | 0.018669 | 9.5023
25 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 46695.0959 | 347.3961 | 0.018669 | 9.5023
26 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 45036.0388 | 347.3009 | 0.019354 | 9.851
27 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 44813.278 | 347.2876 | 0.01945 | 9.8998
28 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 43700 | 347.2191 | 0.019943 | 10.151
29 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 43060.8365 | 347.1781 | 0.020238 | 10.3011
30 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 42140.1465 | 347.117 | 0.020678 | 10.5252
31 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 42140.1465 | 347.117 | 0.020678 | 10.5252
32 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 40405.9041 | 346.9943 | 0.021562 | 10.975
33 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 39685.5161 | 346.9402 | 0.021952 | 11.1733
34 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 39157.6824 | 346.8993 | 0.022246 | 11.3233
35 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 38143.73 | 346.8177 | 0.022835 | 11.6229
36 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 37655.8603 | 346.7768 | 0.023129 | 11.7728
37 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 37655.8603 | 346.7768 | 0.023129 | 11.7728
38 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 36949.8992 | 346.7158 | 0.023569 | 11.9967
39 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 35750.2703 | 346.6067 | 0.024356 | 12.3973
40 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 35609.7561 | 346.5934 | 0.024452 | 12.4459
41 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 35049.0506 | 346.5395 | 0.024841 | 12.6441
42 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 34636.7054 | 346.4987 | 0.025135 | 12.7938
43 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 34636.7054 | 346.4987 | 0.025135 | 12.7938
44 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 33840.9912 | 346.4172 | 0.025723 | 13.0931
45 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 33456.4254 | 346.3764 | 0.026017 | 13.2428
46 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 33017.8682 | 346.3288 | 0.026361 | 13.4178
47 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 32897.9762 | 346.3156 | 0.026457 | 13.4664
48 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 31943.6261 | 346.2067 | 0.027243 | 13.8666
49 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 31831.3953 | 346.1935 | 0.027338 | 13.9152
50 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 31382.6144 | 346.1396 | 0.027727 | 14.1131
51 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 31051.6194 | 346.0989 | 0.028021 | 14.2627
52 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 30410.5776 | 346.0176 | 0.028608 | 14.5616
53 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 30410.5776 | 346.0176 | 0.028608 | 14.5616
54 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 30099.6678 | 345.977 | 0.028902 | 14.7111
55 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 30000 | 345.9638 | 0.028998 | 14.7597
56 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 29744.2312 | 345.9295 | 0.029245 | 14.8859
57 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 29646.8996 | 345.9162 | 0.029341 | 14.9345
58 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 29155.5217 | 345.8482 | 0.029832 | 15.1847
59 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 28869.6244 | 345.8076 | 0.030126 | 15.3342
60 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 28410.6164 | 345.7407 | 0.03061 | 15.5804
61 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 28410.6164 | 345.7407 | 0.03061 | 15.5804
62 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 28139.0734 | 345.7001 | 0.030903 | 15.7298
63 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 28051.9481 | 345.6869 | 0.030999 | 15.7784
64 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 27611.6259 | 345.619 | 0.03149 | 16.0284
65 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 27355.0725 | 345.5784 | 0.031783 | 16.1778
66 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 27312.2294 | 345.5716 | 0.031833 | 16.203
67 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 27272.7273 | 345.5652 | 0.031879 | 16.2263
68 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 27061.1841 | 345.531 | 0.032126 | 16.3523
69 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 26573.0208 | 345.45 | 0.032713 | 16.6508
70 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 26335.3216 | 345.4095 | 0.033006 | 16.8001
71 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 26335.3216 | 345.4095 | 0.033006 | 16.8001
72 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 25952.83 | 345.3427 | 0.033489 | 17.046
73 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 25726.0494 | 345.3022 | 0.033782 | 17.1952
74 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 25653.2067 | 345.289 | 0.033878 | 17.2437
75 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 25355.2354 | 345.2344 | 0.034273 | 17.445
76 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 25284.4744 | 345.2213 | 0.034368 | 17.4935
77 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 25033.1887 | 345.174 | 0.034711 | 17.6679
78 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 25000 | 345.1677 | 0.034757 | 17.6912
79 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 24822.13 | 345.1335 | 0.035004 | 17.817
80 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 24822.13 | 345.1335 | 0.035004 | 17.817
81 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 24410.7924 | 345.0527 | 0.03559 | 18.1151
82 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 24210.0562 | 345.0122 | 0.035883 | 18.2643
83 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 23949.571 | 344.9588 | 0.03627 | 18.4615
84 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 23886.4286 | 344.9456 | 0.036365 | 18.5099
85 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 23632.3851 | 344.8921 | 0.036753 | 18.7074
86 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 23379.2783 | 344.8376 | 0.037148 | 18.9085
87 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 23319.1035 | 344.8245 | 0.037243 | 18.9569
88 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 23105.1993 | 344.7773 | 0.037586 | 19.1311
89 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 23076.9231 | 344.771 | 0.037631 | 19.1543
90 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 23076.9231 | 344.771 | 0.037631 | 19.1543
91 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 22925.2822 | 344.7369 | 0.037878 | 19.2801
92 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 22573.9646 | 344.6562 | 0.038463 | 19.5778
93 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 22573.9646 | 344.6562 | 0.038463 | 19.5778
94 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 22402.1949 | 344.6159 | 0.038756 | 19.7268
95 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 22346.9388 | 344.6028 | 0.038851 | 19.7752
96 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 22178.9807 | 344.5626 | 0.039143 | 19.9237
97 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 22124.8189 | 344.5495 | 0.039238 | 19.9721
98 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 21906.6937 | 344.496 | 0.039626 | 20.1694
99 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 21906.6937 | 344.496 | 0.039626 | 20.1694
100 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 21689.032 | 344.4417 | 0.04002 | 20.3702
101 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 21452.9504 | 344.3815 | 0.040457 | 20.5926
102 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 21452.9504 | 344.3815 | 0.040457 | 20.5926
103 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 21428.5714 | 344.3752 | 0.040503 | 20.6158
104 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 21297.7585 | 344.3412 | 0.040749 | 20.7414
105 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 21247.8103 | 344.3281 | 0.040844 | 20.7898
106 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 20994.2218 | 344.2607 | 0.041334 | 21.0388
107 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 20845.5729 | 344.2205 | 0.041626 | 21.1876
108 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 20797.7208 | 344.2074 | 0.041721 | 21.2359
109 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 20652.1673 | 344.1673 | 0.042012 | 21.3843
110 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 20415.879 | 344.1009 | 0.042495 | 21.6297
111 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 20226.7064 | 344.0467 | 0.042889 | 21.8303
112 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 20226.7064 | 344.0467 | 0.042889 | 21.8303
113 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 20044.2478 | 343.9934 | 0.043276 | 22.0273
114 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 20021.2351 | 343.9866 | 0.043325 | 22.0524
115 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 20000 | 343.9803 | 0.043371 | 22.0756
116 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 19886.001 | 343.9464 | 0.043617 | 22.201
117 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 19842.4483 | 343.9334 | 0.043712 | 22.2493
118 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 19663.7067 | 343.8792 | 0.044106 | 22.4498
119 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 19621.1213 | 343.8661 | 0.044201 | 22.4981
120 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 19449.3783 | 343.8129 | 0.044587 | 22.695
121 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 19322.0281 | 343.7729 | 0.044879 | 22.8432
122 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 19322.0281 | 343.7729 | 0.044879 | 22.8432
123 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 19155.459 | 343.7197 | 0.045265 | 23.0401
124 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 19115.0442 | 343.7067 | 0.04536 | 23.0883
125 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 18949.1126 | 343.6526 | 0.045754 | 23.2887
126 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 18807.6243 | 343.6057 | 0.046095 | 23.4622
127 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 18788.8843 | 343.5994 | 0.04614 | 23.4854
128 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 18768.6625 | 343.5926 | 0.04619 | 23.5105
129 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 18750 | 343.5864 | 0.046235 | 23.5337
130 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 18611.4587 | 343.5395 | 0.046576 | 23.7072
131 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 18454.1188 | 343.4855 | 0.046969 | 23.9075
132 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 18416.6064 | 343.4724 | 0.047064 | 23.9557
133 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 18265.221 | 343.4194 | 0.047451 | 24.1524
134 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 18265.221 | 343.4194 | 0.047451 | 24.1524
135 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 18152.8612 | 343.3794 | 0.047742 | 24.3004
136 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 18005.7634 | 343.3263 | 0.048128 | 24.497
137 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 17970.0499 | 343.3133 | 0.048223 | 24.5453
138 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 17823.3252 | 343.2594 | 0.048616 | 24.7454
139 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 17788.3311 | 343.2463 | 0.04871 | 24.7936
140 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 17698.0939 | 343.2126 | 0.048956 | 24.9187
141 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 17681.4988 | 343.2063 | 0.049002 | 24.9419
142 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 17663.5893 | 343.1996 | 0.049051 | 24.9669
143 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 17524.2837 | 343.1466 | 0.049437 | 25.1634
144 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 17524.2837 | 343.1466 | 0.049437 | 25.1634
145 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 17384.7198 | 343.0926 | 0.04983 | 25.3634
146 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 17216.9811 | 343.0267 | 0.050311 | 25.6081
147 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 17117.1126 | 342.9868 | 0.050601 | 25.756
148 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 17084.8341 | 342.9738 | 0.050696 | 25.8041
149 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 16986.2614 | 342.9339 | 0.050987 | 25.9523
150 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 16823.8048 | 342.867 | 0.051474 | 26.2004
151 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 16792.622 | 342.8541 | 0.051569 | 26.2485
152 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 16712.1814 | 342.8204 | 0.051814 | 26.3736
153 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 16697.383 | 342.8141 | 0.05186 | 26.3967
154 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 16697.383 | 342.8141 | 0.05186 | 26.3967
155 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 16557.112 | 342.7545 | 0.052295 | 26.618
156 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 16432.4733 | 342.7007 | 0.052687 | 26.8177
157 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 16432.4733 | 342.7007 | 0.052687 | 26.8177
158 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 16311.8433 | 342.6479 | 0.053073 | 27.014
159 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 16282.5279 | 342.6349 | 0.053167 | 27.0621
160 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 16193.1778 | 342.5951 | 0.053457 | 27.2098
161 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 16164.2869 | 342.5822 | 0.053552 | 27.2579
162 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 16076.023 | 342.5423 | 0.053843 | 27.406
163 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 16076.023 | 342.5423 | 0.053843 | 27.406
164 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 15902.4745 | 342.4627 | 0.054424 | 27.7018
165 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 15830.3176 | 342.4291 | 0.054669 | 27.8267
166 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 15830.3176 | 342.4291 | 0.054669 | 27.8267
167 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 15817.0391 | 342.4229 | 0.054715 | 27.8498
168 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 15718.3363 | 342.3763 | 0.055055 | 28.0228
169 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 15691.1135 | 342.3634 | 0.055149 | 28.0708
170 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 15579.128 | 342.3097 | 0.055541 | 28.2704
171 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 15470.6601 | 342.257 | 0.055926 | 28.4664
172 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 15444.2877 | 342.244 | 0.05602 | 28.5144
173 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 15337.8679 | 342.1914 | 0.056405 | 28.71
174 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 15258.3763 | 342.1517 | 0.056695 | 28.8579
175 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 15101.9468 | 342.0722 | 0.057276 | 29.1534
176 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 15101.9468 | 342.0722 | 0.057276 | 29.1534
177 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 15036.8569 | 342.0387 | 0.057521 | 29.2781
178 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 15024.8756 | 342.0325 | 0.057566 | 29.3012
179 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 14935.784 | 341.986 | 0.057906 | 29.474
180 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 14911.2023 | 341.9731 | 0.058 | 29.522
181 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 14810.0365 | 341.9195 | 0.058392 | 29.7213
182 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 14785.8666 | 341.9066 | 0.058486 | 29.7693
183 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 14711.9799 | 341.8669 | 0.058776 | 29.9171
184 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 14591.8421 | 341.8015 | 0.059254 | 30.1605
185 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 14591.8421 | 341.8015 | 0.059254 | 30.1605
186 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 14519.8771 | 341.7619 | 0.059545 | 30.3082
187 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 14378.1531 | 341.6826 | 0.060124 | 30.6033
188 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 14319.1406 | 341.6491 | 0.060369 | 30.7279
189 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 14308.2754 | 341.6429 | 0.060415 | 30.751
190 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 14296.5453 | 341.6363 | 0.060463 | 30.7759
191 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 14227.4566 | 341.5966 | 0.060754 | 30.9235
192 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 14113.3078 | 341.5303 | 0.061239 | 31.1706
193 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 14091.3568 | 341.5174 | 0.061333 | 31.2186
194 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 14024.2325 | 341.4778 | 0.061623 | 31.3661
195 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 14024.2325 | 341.4778 | 0.061623 | 31.3661
196 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 13915.0227 | 341.4126 | 0.062101 | 31.6092
197 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 13849.5639 | 341.373 | 0.062391 | 31.7568
198 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 13741.3752 | 341.3067 | 0.062876 | 32.0037
199 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 13720.5651 | 341.2939 | 0.06297 | 32.0516
200 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 13666.8171 | 341.2605 | 0.063214 | 32.1761
201 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 13646.232 | 341.2476 | 0.063308 | 32.224
202 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 13583.2719 | 341.2081 | 0.063598 | 32.3714
203 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 13583.2719 | 341.2081 | 0.063598 | 32.3714
204 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 13459.1639 | 341.1291 | 0.064177 | 32.6661
205 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 13397.9143 | 341.0896 | 0.064467 | 32.8135
206 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 13317.7543 | 341.0373 | 0.06485 | 33.0085
207 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 13298.2066 | 341.0245 | 0.064944 | 33.0563
208 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 13139.5248 | 340.9189 | 0.065718 | 33.4503
209 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 13120.4964 | 340.9061 | 0.065812 | 33.4982
210 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 13052.5069 | 340.8599 | 0.06615 | 33.6704
211 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 12994.8946 | 340.8205 | 0.066439 | 33.8177
212 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 12881.2606 | 340.7416 | 0.067018 | 34.112
213 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 12881.2606 | 340.7416 | 0.067018 | 34.112
214 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 12825.1469 | 340.7022 | 0.067307 | 34.2592
215 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 12751.6753 | 340.6501 | 0.067689 | 34.4539
216 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 12733.753 | 340.6372 | 0.067783 | 34.5018
217 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 12588.1823 | 340.5319 | 0.068557 | 34.8953
218 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 12508.2917 | 340.4731 | 0.068988 | 35.1151
219 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 12508.2917 | 340.4731 | 0.068988 | 35.1151
220 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 12455.3737 | 340.4337 | 0.069277 | 35.2622
221 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 12350.9416 | 340.355 | 0.069855 | 35.5562
222 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 12299.3439 | 340.3157 | 0.070144 | 35.7033
223 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 12231.7573 | 340.2637 | 0.070526 | 35.8978
224 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 12081.2459 | 340.1458 | 0.071392 | 36.3386
225 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 12081.2459 | 340.1458 | 0.071392 | 36.3386
226 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 12007.6414 | 340.0871 | 0.071824 | 36.5582
227 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 11958.8666 | 340.0478 | 0.072112 | 36.7051
228 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 11878.1146 | 339.9821 | 0.072595 | 36.951
229 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 11862.5622 | 339.9693 | 0.072689 | 36.9988
230 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 11752.5752 | 339.8782 | 0.073359 | 37.3399
231 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 11752.5752 | 339.8782 | 0.073359 | 37.3399
232 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 11613.5584 | 339.7606 | 0.074225 | 37.7803
233 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 11560.258 | 339.7147 | 0.074562 | 37.9519
234 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 11545.5263 | 339.702 | 0.074655 | 37.9996
235 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 11425.7276 | 339.5972 | 0.075426 | 38.392
236 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 11411.3366 | 339.5845 | 0.07552 | 38.4397
237 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 11309.5218 | 339.4936 | 0.07619 | 38.7805
238 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 11180.7315 | 339.3762 | 0.077054 | 39.2203
239 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 11131.3215 | 339.3305 | 0.07739 | 39.3917
240 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 11117.6621 | 339.3178 | 0.077484 | 39.4393
241 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 11006.5354 | 339.2133 | 0.078254 | 39.8313
242 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 10898.6597 | 339.1098 | 0.079016 | 40.2193
243 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 10779.0075 | 338.9927 | 0.079879 | 40.6586
244 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 10733.077 | 338.9471 | 0.080216 | 40.8298
245 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 10617.0136 | 338.8301 | 0.081078 | 41.2689
246 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 10617.0136 | 338.8301 | 0.081078 | 41.2689
247 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 10516.6035 | 338.7269 | 0.08184 | 41.6565
248 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 10362.3443 | 338.5646 | 0.083038 | 42.2663
249 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 10362.3443 | 338.5646 | 0.083038 | 42.2663
250 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 10254.1198 | 338.4479 | 0.0839 | 42.7049
251 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 10016.3674 | 338.1829 | 0.085857 | 43.7012
252 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 9915.2138 | 338.0665 | 0.086718 | 44.1393
253 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 9692.747 | 337.8021 | 0.088673 | 45.1344
254 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 9597.9933 | 337.686 | 0.089532 | 45.572
255 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 9389.384 | 337.4222 | 0.091485 | 46.566

## Sheet2 (hidden)
Excitation | IAA100 Gain and Shunt Settings V6, 5/11/2021
5
10